Proposal to the ISCD Position Development Conference (PDC) at the ISCD 2022 Annual Meeting

## Topic Information and Rationale:

Add a “Central Radiofrequency Echographic Multi Spectrometry (REMS) for Diagnosis” and “REMS Fragility Score (FS)” section to the “Indications for Bone Mineral Density (BMD) Testing” issued as part of the “Adult Official Positions of the ISCD as updated in 2019”

Proposal compiled by: Andrew Bush, Curtis Foy, Yevgeniya Kushchayeva, Ed McLaughlin and Kimberly Zambito

Contact Information: 		    Andrew P. Bush, MD, FAAOS, CWSP

Central Carolina Orthopaedic Associates, PLLC

1139 Carthage Street, Suite 101

Sanford, NC 27330

919.774.1355

abush@centralcarolinaortho.com

INTRODUCTION:

Radiofrequency Echographic Multi Spectrometry (REMS) is a quantitative ultrasound technology that has been developed to measure, assess and report specific and clinically significant characteristics of human bone in a manner that complies with the WHO established criteria for the diagnosis of osteoporosis (1). The predecessor of the REMS technology, quantitative ultrasound (QUS) (for the calcaneus) was found to have clinical relevance as a predictor of fracture risk (2,3). However, due to the significant variability in the sensitivity and specificity of the QUS in determining the bone mineral density (BMD), it could not be used as a clinically valid method to diagnose and treat osteoporosis based on the WHO guidelines (1-3). In spite of the fact that QUS was universally recognized as a powerful tool for fracture prediction (2,4-6), its clinical use was deemed to be too limited and it was relegated to be used as an adjuvant screening tool (3).

REMS technology is a new generation of ultrasound systems that has overcome the limitations of QUS in several important ways. Advancements in ultrasound technology have resulted in the development of innovative hardware that now allows for the assessment of axial bones with REMS. REMS is not limited only to peripheral site evaluations (i.e., calcaneus, tibia, phalanx), which was the case with QUS (3). The development of REMS made possible to sonographically assess the lumbar spine and proximal femurs – anatomical regions that up to now, could only be measured by dual-energy x-ray absorptiometry (DXA). In addition, innovative and proprietary software development has allowed for the acquisition and automatic analysis of backscattered ultrasound waves, not only to generate B-mode images for clinical use during the REMS assessment, but also to assess bone density as well as bone quality properties. The results are rapidly reported in real-time and therefore are immediately available to both the examiner and patient.

WHAT IS RADIOFREQUENCY ECHOGRAPHIC MULTI SPECTROMETRY (REMS):

REMS is a novel technology that propagates sonographic waves (ultrasound) through the bone at a specific frequency in a targeted region of the human body and receives the backscattered radiofrequency signals (RF) after they have interacted with the bone microarchitecture (7,8). Multiple sonographic signals are acquired from these specific regions of interest (ROI) allowing for both conventional B-mode echographic images and unfiltered (raw) RF. The RF signals are then filtered, amplified, digitized and stored for analysis (7). The underlying concept for the collection and processing of raw RF signals is based on the fact that the RF signals that have interacted with the bone contain information about the bone status, including both quantity and quality (3,7-9).

REMS scanning relies on an echographic portable device, EchoStation (EchoS) manufactured by Echolight S.p.A, Lecce Italy. EchoS consists of a dedicated ultrasound unit which is connected to a convex echographic hand-held transducer (probe) that is used to perform the acquisition. The ultrasound unit is connected by USB cable to a PC hard-disk drive (19” Panel PC or laptop computer), where the acquired data is stored and the analysis is performed. This construct is mounted on a mobile cart (10).

Scans are performed with a hand-held convex transducer that emits an operating frequency of 3.5 MHz. To perform the assessment the examiner will visualize the lumbar spine by placing the transducer over the unclothed abdomen of the patient being examined. For each scan, the examiner will set the transducer focus and scan depth parameters for the acquisition in order to detect the sought bone interfaces (7,8).

Utilizing knowledge of the anatomy of the lumbar spine, L3 is identified as the most prominent vertebral body. Transducer focus and scan depth are set and L4, L2 and L1 are then identified in their relative positions with respect to L3. Data acquisition begins at L1 and the examiner is prompted to move distal in a sequential manner until the scan is completed up to L4. Total acquisition time is approximately 80 seconds. During the data processing that lasts 1-2 minutes, the acquired RF data run through multiple elaboration steps. Segmentation algorithms are then applied to automatically identify the targeted bone interfaces within the sequence of the frames of acquired echographic images; such algorithms check for the correct position of the detected interface in the frame, excluding those whose vertebral interfaces fall outside the expected range (7). If the “minimal frame condition” is not satisfied due to insufficient number of acceptable frames, the data set will be considered as “noisy” and the scan will not be accepted. In this case, a report will not be generated and the examiner will be asked to repeat the examination. If the data acquisition was considered to be sufficient, calculations of bone quantity (BMD determination) and quality (microarchitecture assessment) will be performed and reported on each detected ROI.

Examinations for the hip follow the same sequence (8). However, the transducer is held in a stationary position over the proximal femur to allow for the scan to start. Data processing follows the same steps as previously described.

The bone quantity measurement is expressed as the BMD value. It is determined through the processing of specific backscattered RF signals (as described in the previous paragraph) that results in the calculation of an Osteoporosis Score (O.S.) (7,8). The obtained O.S. value of the patient is then statistically analyzed (by Pearson correlation coefficient) in comparison to a proprietary database of reference ultrasound spectral models classified as “normal”, “osteopenic” (low bone density) or “osteoporotic”. The O.S. is used to directly calculate the BMD using a series of linear equations. The corresponding T-score and Z-score values are then derived from a normative database (National Health and Nutrition Examination Survey, NHANES). This approach has been validated across multiple clinical studies (12-17).

The bone quality measurement is determined in a similar manner. The appropriate RF signal is acquired and upon completion of the standard rigorous analytical processes, the values of the RF signal segments from the selected ROI are statistically compared (by Pearson correlation coefficient) to “frail” (fractured) and “non-frail” (not fractured) spectral signals from an established proprietary database. The Fragility Score (FS) is derived by computing the averaged sum of all the RF segments from the selected ROI. The FS has shown the ability to discriminate between patients who experienced an osteoporotic fracture from those who did not (sensitivity = 76%; specificity = 68%) and compared to DXA BMD (73% and 66% respectively) (10).

CAN REMS BE CONFIDENTLY USED TO DIAGNOSE OSTEOPOROSIS ACCORDING TO THE WHO CRITERIA:

REMS made possible axial sonographic assessments - the lumbar spine and proximal femurs – regions that could only, up to now, be measured by DXA. Because REMS displays this capability, its performance compared to DXA, which is well recognized as the standard method for BMD determination based on the WHO criteria (3). REMS has shown the capacity to provide a reliable (accurate) and reproducible (precise) measurement of BMD as compared to DXA in multiple clinical studies (12-17). In the cited studies, multiple authors converged on the conclusion that there is a very high correlation between DXA-measured and REMS-measured BMD and T-score values in the assessed populations, and that REMS had high accuracy, sensitivity and specificity for the diagnosis of osteoporosis compared to DXA. Even though REMS is a new technology in the United States, REMS technology has been fully vetted and cleared for clinical use to measure BMD and bone quality in the United States following the regular and rigorous Food and Drug Administration (FDA) approval process (18). During the last years, REMS has been increasingly used in Europe thanks to wide clinical acceptance within the European Union (11-13,19). Recently, REMS technology has been recognized as “the first ultrasound-based method to diagnose osteoporosis and assess the risk of fragility fractures” by the Italian National Institute of Health and thus it has been officially included in the Italian Ministerial Guidelines for the “diagnosis, risk stratification and continuity of care following fragility fractures” (20).

CAN REMS BE CONFIDENTLY USED TO MONITOR OSTEOPOROSIS TREATMENT ACCORDING TO THE WHO CRITERIA:

It has been established that REMS is comparable to DXA in determining BMD and generating T-scores and Z-scores that are clinically significant and valid in diagnosing osteoporosis. Several reference parameters have been reported (12,13) that document a high accuracy, reproducibility and precision of the REMS technology:

| Quantitative Assessment (12) | Vertebra | Femoral Neck |
| --- | --- | --- |
| Smallest Detectable Difference (SDD) [g/cm2] | 0.009 | 0.006 |
| Intra-operator Repeatability (RMS-CV) [%] | 0.38% | 0.32% |
| Inter-operator Repeatability (RMS-CV) [%] | 0.54% | 0.48% |
| Diagnostic Agreement with DXA | 88.8% | 88.2% |

(Abbreviations: RMS - root mean square; CV - coefficient of variation.)

Based on the ISCD Official Position (21) the minimum acceptable precision for an individual DXA technologist is:

| Region Of Interest | Precision Error | Least Significant Change |
| --- | --- | --- |
| Lumbar Spine | 1.9 % | 5.3 % |
| Total Hip | 1.8 % | 5.0 % |
| Femoral Neck | 2.5 % | 6.9 % |

As shown in the European studies, REMS accuracy and precision was validated with respect to DXA, demonstrating its effective clinical use for the diagnosis of osteoporosis according to the WHO criteria (13). Therefore, as documented in the above charts, REMS can be used to monitor BMD and provide information that is accurate, precise and reproducible in a patient over time without many of the concerns that plague DXA.

CAN REMS BE CONFIDENTLY USED TO MONITOR OSTEOPOROSIS IN PATIENTS WITH METABOLIC DISEASE CONDITIONS THAT AFFECT BMD:

A recent study demonstrated the unique ability of REMS to provide more accurate BMD measurements in patients where DXA evaluations are recognized to be potentially flawed, such as for the assessment of osteoporosis in women with type 2 diabetes (16). Also, several studies have documented the ability of REMS, based on the analysis of unfiltered backscattered RF waves, to eliminate RF that are not associated with bone mass, therefore automatically eliminating artifacts (15). Consequently, while artifacts do not have an effect on REMS-derived BMD results, they can significantly impact the obtained DXA outcome (22, 23).

CAN THE REMS-DERIVED FRAGILITY SCORE DETERMINE FRACTURE RISK:

The single most compelling aspect of REMS technology is the ability to assess bone quality and provide a predictive tool for fracture risk utilizing the FS parameter. It is well documented in the literature that the backscatter RF waves emanating from the insonification of the target bone contain information on the microarchitecture of the bone (2-10,12,24). From basic science experimentation conducing spectral analysis on bone biopsy specimens, it is suggested that factors other than BMD, such as bone crystal size and collagen maturity, have a predictive value for fracture risk (25,26). Although the techniques described in the mentioned articles do not recommend the use of bone biopsies as a clinically valid method to obtain this information, they overall support the evidence that multiple factors determine bone strength. The authors however, concluded that other non-invasive methods should be further investigated to obtain information on bone microarchitecture.

Other types of non-invasive modalities have been studied as clinically available methods to determine fracture risk. Quantitative Computed Tomography (QCT) has been shown to provide information on bone microstructure (9,11,27) and bone strength determined by finite-element analysis (FEA) (28). A 5-year prospective study found a statistically significant correlation between low femoral strength determined by QCT and incident hip fractures (3). However, the costs, radiation exposure and limited access make this method limited in its clinical usefulness. High-resolution peripheral QCT (HR-pQCT) also has limited clinical usefulness at this time (29). Magnetic resonance imaging has similar costs and availability constraints as QCT and therefore results an unlikely method to determine bone strength (30).

Trabecular bone score (TBS), is a mathematically calculated reconstruction of the three-dimensional microarchitecture of bone derived by processing the pixel gray level variations found in two-dimensional DXA images (31). It has a significant potential for clinical applications – TBS seems to be related to trabecular microarchitecture and has a better ability to predict fractures over DXA-derived BMD in certain clinical situations. However, it has certain limitations based on availability – not all DXA machines have the necessary software. TBS also seems to be as susceptible to artifacts as DXA.

A few studies investigated the ability of an advanced ultrasound-based approach (apparent integrated backscatter, AIB) to evaluate the microarchitecture of both portioned femoral head specimens as well as intact femoral heads (9,27). By comparing the results obtained from the analysis of backscattered RF waves (AIB) from bone specimens with those obtained by QCT, a strong correlation between AIB and bone trabecular bone volume fraction (BV/TV) could be found.

Sustaining fractures has a considerable impact on the quality of life, thus making prevention and monitoring in response to treatments of outmost importance in frail patients (32). A cross-sectional clinical study examined the ability of the REMS derived FS to distinguish between fractured and non-fractured patients (10). The FS has shown the ability to discriminate between patients who experienced an osteoporotic fracture from those who did not (sensitivity = 76%; specificity = 68%) as compared to DXA BMD (73% and 66% respectively) (10).  A recent work resumed a 5-year longitudinal study validating the use of the REMS-derived FS as a predictive tool for fracture risk (33).

CAN REMS BE USED IN DETERMINING BONE QUALITY DURING PREGNANCY AND IN CHILDREN:

REMS has been successfully used to measure BMD changes during pregnancy. The study conducted at the unit of Obstetrics and Gynecology, University of Parma, enrolled pregnant women from February to December 2020 (40). In this exploratory and proof of concept study, a BMD reduction was found in pregnant women, demonstrating that measurable changes in BMD can be detected by means of the REMS technology (40). Additional preliminary results from another study documented the successful use of REMS to measure bilateral femur BMD in a group of pregnant women (41). These studies would be unfeasible using DXA due to the ionizing radiation exposure for both the participants and their unborn children (40,41).

Pediatric BMD assessment is currently performed by DXA as a “whole-body” scan. Results have been deemed to be reliable but only if appropriately performed. Pediatric DXA examinations have the same limitations as adult DXA examinations where the results depend on correct patient positioning, consistency of scanning technique and facility for LSC monitoring, correct ROI settings, difference of scanning system and the elimination of artifacts (42). Pediatric DXA examinations are not routinely done due to concerns related to the use of ionizing radiation on a growing and developing body.

As opposed to DXA, REMS technology is safe for pediatric use due to the employment of non-ionizing energy source. Also, there are fewer requirements that the examiner performing a REMS evaluation has to be concerned with as compared to the DXA. Indeed, BMD evaluation using REMS is independent of patient positioning and artifacts in addition to the capacity of automatic determination of ROI for the diagnostic evaluation.

Current REMS algorithms have a low age limit of 20 years old. Data compilation for the development of statistical tables that would be required to analyze younger children were delayed by the recent and ongoing COVID-19 pandemic. In the meanwhile, clinical studies and validation of the database for analyses is ongoing with the goal to promote the use of BMD for routine screenings of the pediatric population.

CAN REMS BE USED IN THE SHORT-TERM MONITORING OF PHARMACEUTICAL OSTEOPOROSIS TREATMENT:

REMS technology has demonstrated high performance in terms of accuracy and precision of BMD measurement. Its low LSC value allows for clinically meaningful densitometric studies at six-month intervals, unlike DXA where the interval for clinically meaningful studies has been established at one to two years. Therefore, it is anticipated that REMS technology will be used to monitor the short-term effectiveness of pharmaceutical treatments of osteoporosis. A preliminary study performed in collaboration between the Institute of Clinical Physiology and the departments of Rheumatology and Oncology at two hospitals in Lecce, in Italy, concluded that REMS technology displays high precision and repeatability performances allowing for serial BMD measurements of the femoral neck at a period of six months in a group of patients under Denosumab treatment (39,43,44). Additional studies are under assessment.

CAN REMS DETERMINE BODY COMPOSITION AND DIAGNOSE SARCOPENIA:

The body composition module uses the same data acquired during the REMS scan of lumbar vertebrae, in order to determine the Body Fat expressed in percentage and the Basal Metabolic Rate (BMR). The process takes a few seconds and the resulting parameters are based on a multiple regression model that accounts for abdominal measurements of total thickness of soft tissues between skin and lumbar vertebrae, subcutaneous fat thickness and muscle thickness and then combines them with patient characteristics (age, sex, height and weight).

The measurement of sarcopenia is a parameter that is not available in the present EchoS software package, but is currently under development.

CONFLICTING LITERATURE:

An important part of this proposal for the ISCD PDC committee was to elaborate a comprehensive bibliography of the available peer-reviewed literature to assist in the processing of this proposal for a Policy Statement on REMS. What we found to our surprise, was a plethora of articles and abstracts supporting the use of REMS technology, and the absence of literature discrediting its clinical use and validity. We understand that REMS is still a relatively new technology. However, it has widespread use in Italy and the European Union and it was our expectations that there would be a balance between opposing medical opinions and a healthy debate as to the relevance of REMS in BMD assessment.  We were only able to find a single abstract that questioned clinical validity of REMS technology (47). However, this study has a very limited patient population, has not been completed and there may have been issues related to the insufficient training of the operators performing the REMS scans.

CAN THE ECHOS SYSTEM PROVIDE A REPORT THAT MEETS ISCD STANDARDS:

The generated reports following the completion of a REMS assessment, include not only the standard reporting of BMD, T-sores, Z-scores, diagnosis and ROIs examined such as the lumbar spine (L1-L4) and the bilateral femoral necks and total hips, but also the fragility score (FS), body composition, reference range for fracture risk prediction and B-mode doppler images of the measured ROIs (not for diagnostic purposes). These reports are standardized and are uniform between all REMS systems. Examples of the reports are provided herein:

TECHNICAL ASPECTS OF THE ECHOS SYSTEM:

- Quality control (Q/A Q/C) – the Quality Check module is an additional tool that can be used by the operator every six months to verify that the EchoS is operating at a factory-set performance level. The Quality Check is launched by the operator who is then instructed to follow a sequence of steps. Once the sequence is initiated the operator is instructed to insert the transducer into a Quality Check phantom – a graph is then generated that indicates that the system is performing at the pre-set parameters. If the graph falls out of the expected range, then operator is required to contact Echolight support at: support@echolight.it (48).

- Calibration – there is no need to calibrate the EchoS system. Due to the physics of the components that make up the transducer (piezo electric crystals), there is no degradation of quality or expected failure. If a crystal element were to fail, the failure would be observable on the B-mode image – if enough crystals were corrupted then the automated system would deem the study as “noisy” and a report would not be generated (7). A faulty transducer component, or any component that could experience a malfunction can be easily replaced.

- Software updates – regular software updates and/or enhancements are routine (49) and similar to any computer or smart-phone update. The updates occur remotely and are scheduled during hours that are selected by the operator.

- Comparison between the same and different operators:

| PARAMETER | REMS (12) | DXA (45,46) | DXA (47) |
| --- | --- | --- | --- |
| Smallest Detectable Difference (SDD) | 0.006 - 0.009 g/cm2 | 0.0400-0.047 g/cm2 | 0.030 g/cm2 |
| Least Significant Change (LSC) | 0.38 – 1.05 % | 4.93 – 5.60 % | 4.07 % |
| Intra-Operator Repeatability | 0.32 – 0.36 % | 1.78 – 2.02 % | 1.47 % |
| Inter-Operator Repeatability | 0.48 – 0.54 % | N.D. | N.D. |

Multicenter clinical studies demonstrated a good level of precision (intra-operator repeatability) and repeatability (inter- operator repeatability) (12). In particular, for each reference site 30 patients were scanned, undergoing two consecutive investigations with repositioning in-between, in line with the ISCD guidelines. Precision was assessed on data acquired on two investigations performed by an experienced operator, whereas repeatability was assessed on data acquired on two investigations performed by two different operators, an experienced one and another who had previously received a short training. () Factors associated to high precision and repeatability are explained by to:

- Automation of the data analysis

- No need for any specific patient positioning due to the use of a hand-held transducer

- No equipment component that undergoes a time dependent degradation process

- Training recommendations – as with any procedure there is a learning curve for new operators. Based on the clinical experience of two of the current users here in the United States (Dr. Kimberly Zambito and Dr. Andrew Bush) proficiency in obtaining acceptable REMS results after a single scan occurs somewhere between scans 30 and 40. Additional studies will help examiners further hone their skills. Initial training for inexperienced operators will consisted of 2-to-3-day course involving didactic classroom lectures and hands-on training sessions using human models. Unlike DXA training and calibration which have potential ethical ramifications due to the exposure of the human models to ionizing radiation for no medical reason, there is no such concern with the REMS training. There are plans for a certification program to document and establish proficiency with EchoS use.

- Hardware failure/degradation – as in any operating system, there is a limited time of functional use (life-expectancy) of any piece of equipment determined by the function of the modulus of failure. Unlike a cathode ray tube that will undergo a time dependent degradation due to normal use, the failure of the piezoelectric crystals in the transducer unit is unlikely with normal use. The possibility of failure of the piezo-electric crystals in the transducer would be considered unusual and would be easily recognized – it would result in an adulterated B-mode image. If the theoretical failure of the transducer resulted in the loss of RF data and the “minimal frame condition” could not be satisfied due to insufficient number of acceptable frames of data, the data set will be considered as “noisy” and the scan will not be accepted (7). In this case, a report will not be generated and the examination would be terminated. Technical support would have to be called and the study rescheduled – but the patient would not be given an incorrect scan. The functional life of the equipment can be prolonged with careful use and maintenance of the equipment. Misuse or abuse of equipment or damage caused by an inadvertent event (such as dropping the transducer) can result in equipment failure, requiring the replacement of the equipment. Individual components can be replaced. Up to this time, there have not been any reports of unexpected equipment failure in any of the currently operational EchoS units.

- Sources of error – due to the automation of the data collection and processing, there is very little chance of error. This is the reason why the results of REMS testing are accurate, precise and reliable. However, as in any technique that requires human input, there is one case where erroneous data can be collected resulting in a report with suboptimal BMD determination. It is important that the operator sets parameters for the acquisition [the transducer focus (21-100 mm) and scan depth (60-210 mm)] correctly in order to visualize the target bone interface centered on the screen with the focal line near-touching the surface of the bone being examined (13). If these parameters are correctly set, the exam will result in a correct BMD measurement. It has been noted that incorrect acquisition positioning, does not affect the Fragility Score (FS) value. The fact that repeatability of measurements remains constant, proves that the device outcome is highly reproducible and reliable.

CAN REMS AND ECHOS HAVE A CLINICAL APPLICATION IN THE DIAGNOSIS AND TREATMENT OF OSTEOPOROSIS AND IN THE EVALUATION OF FRACTURE RISK:

Serving patients and providers in a community private practice setting:

REMS technology is implemented the EchoS ultrasound system which is a fantastic clinical tool. With this clinical tool the current users have been able to create and build solid Bone Health Programs that can serve as the Fracture Liaison Service (FLS) for hospitals and surrounding communities. The ability to provide in office examinations and to interact with patients who can watch as their bones are being assessed in real-time with the opportunity to review results immediately upon completion of the examination, is a significant attribute over DXA generated reports sent over from a local imaging center. Also, since REMS is a non-ionizing and safe diagnostic approach, there are no OSHA or FDA compliance issues. The EchoS is portable, which provides the opportunity to not only utilizing the same unit in multiple offices, but also to be able to offer bone health services to individuals who have mobility issues.

Serving patients and providers in underserved communities:

Information about osteoporosis screening is lacking among uninsured and underinsured populations in the literature (34). Gillespie, et al. reported lower utilization rates amount non-Hispanic black women and women of low socioeconomic status, potentially reflecting attitudes toward risk for developing osteoporosis, access to affordable care, and compliance with physician referrals for screening for osteoporosis.  Their findings were consistent with survey data by United States Department of Health and Human Services (35).

The United States Preventive Service Task Force reported that fracture risks were highest for non-Hispanic white woman, the quantity and quality of data pertaining to other racial and ethnic groups were lacking (35). In a more recent report dated June 2018, the USPSTF included the recommendation of use of quantitative ultrasound to screen for osteoporosis, citing QUS had no exposure to radiation and used portable devices, which could increase access to screening in locations where machines that perform central DXA were not available.  While QUS does not measure BMD and no treatment studies use QUS measurements to define treatment threshold; QUS could not be routinely used to initiate treatment without further DXA measurement (36,37).  Furthermore, the authors stated a method for converting or adapting QUS results to the DXA scale needed to be developed (36,37).

As was described in the above section CAN REMS BE CONFIDENTLY USED TO DIAGNOSE OSTEOPOROSIS ACCORDING TO THE WHO CRITERIA REMS technology and the EchoS have shown the capacity to provide a reliable (accurate) and reproducible (precise) measurement of BMD as compared to DXA in multiple clinical studies (12-17). Therefore, EchoS with REMS technology is the only system that is available in this clinical situation that will determine valid BMD and T-score values allowing a diagnosis to be made and the correct treatment initiated. The accuracy of the REMS will also allow for early monitoring of medication effectiveness. Additional information obtained would also include characterization of fracture risk using the Fragility Score (10, 32,33).

EchoS is an easily transportable system with immense potential to reach multitudes of patients who have fallen victim to the disparities that are present in our healthcare delivery system. These disparities have resulted in unnecessary and preventable patient pain and suffering (34) and unnecessary financial costs due to potentially preventable fragility fractures that were not diagnosed and/or treated in a timely manner (34-37).

Serving patients and providers in the long-term care and skilled nursing setting:

DXA technology has been available since 1987 and while it remains an important tool and considered by many as the “gold standard” for bone density measurement, there are issues that present challenges in deploying it to a wider at-risk and underserved patient population – for example, nursing home and assisted living facility residents. This population is especially at high risk for fracture, yet remains one of the largest patient populations that fail to get screened (and subsequently treated) because of logistical issues such as transportation costs, mobility issues, isolation requirements during the pandemic, etc. Being able to bring the technology to the patients (bedridden subjects also) would immediately improve screening rates. Also, with the BMD and T-score results obtained by the portable EchoS, treatment plans can be designed for these underserved and vulnerable patients with the goal of decreasing their fracture risk.

The physical and financial burdens that fragility fractures impose upon the elderly is significant.  Having an additional portable and accurate tool at our disposal that not only can diagnose but also monitor treated patients, can be a valuable strategy to fight against the burden of osteoporotic fragility fractures.

In addition, even though DXA technology has been established since 1987 and is in widespread use, errors in interpretation continue to be a problem despite best efforts being made to educate those who interpret the studies.  As reimbursement has been unwisely cut over the past two decades, thousands of machines have disappeared across the United States making access to DXA even more challenging.  Having a method of bone health assessment that can predictably and reliably detect significantly smaller changes in bone density, that is portable and does not produce ionizing radiation, can help clinicians make treatment decisions with greater confidence.

Serving patients and providers /researchers in the larger university and medical center setting:

REMS technology has significant benefits in actual clinical practice. REMS studies can be easily performed in the office setting with immediate reporting to the provider and patient. This strategy, not only can increase the number of individuals screened and treated for osteoporosis, but it also can be time-efficient and may decrease the number of visits for the patients. In addition, the portable REMS device can be a valuable technique for preoperative evaluation of patients prior to spinal or orthopedic surgeries. Radiation-free REMS technology is the first diagnostic modality that can be used during pregnancy to follow bone loss and help in the preparation for immediate treatment after delivery.

REMS has an important and unique value in the evaluation of patients with skeletal deformities and/or presence of instrumentation in comparison with other available modalities where the hardware will adversely affect the results of the study (i.e., Osteogenesis Imperfecta). In these cases, REMS technology is a critical tool for diagnosis of bone health status, disease progression and monitoring of treatment effect. Furthermore, insurance approvals for the appropriate osteoporosis therapy are often based on BMD measurements that cannot be reliably received in many patients including ones with significant bone deformities or presence of instrumentation. REMS can be successfully used to obtain the BMD to justify medical necessity in bone-targeted treatment.

REMS is a ground-breaking tool that can reawake osteoporosis awareness and increase osteoporosis screening due to a high-quality, reliable US-based and readily accessible technology.

SUMMARY OF HOW REMS TECHNOLOGY AND THE ECHOS CAN BENEFIT AND CONTRIBUTE VALUE TO CURRENT BONE HEALTH CARE PRACTICES:

Some of the areas where REMS has proven to be a valuable addition to the established modalities for the bone health assessment are listed below:

- REMS can be easily accessible for:

  - people in nursing facilities, rehabilitation centers, and any other place with no available DXA scanner such as small villages and towns.

  - People who are immobile (after stroke, with muscular dystrophies, recent bone fracture, etc.). Evaluation of bone health with REMS can be done even at home and on bedridden subjects.

- REMS can be used in the cohorts where DXA study cannot be completed such as:

  - During pregnancy

  - Presence of instrumentation in the ROIs (i.e., Osteogenesis Imperfecta, fractures, presence of metallic body, presence of orthopedic surgery outcomes; etc)

  - Presence of severe bone deformities (scoliosis, rotoscoliosis; etc.)

  - Presence of significant bone degenerative changes (arthritis, etc.), peripheral/vascular calcifications, implants, etc.

- Critical aspect that REMS technology does not require to be performed in Radiology departments will lead to a number of benefits:

  - Decrease load on Radiology Departments

  - Will make reporting more unified since technique is not technologist/interpreter dependent

  - Will eliminate an additional appointment for the bone imaging that is critical for patients who have limited mobility and require assistance

  - The REMS can be done during/before the appointment with the primary care provider or any specialist who focused on bone health

  - REMS medical report will be immediately available for review by the provider, avoiding the need for several medical consultations as well as reducing the waiting lists

  - No issues related to the disposal of radiologic reports

- Expected/Predicted benefits:

  - Availability REMS in offices/clinics will:

    - significantly decrease the time before treatment can be initiated,

    - more people will receive osteoporosis screening and initiate appropriate treatment

    - less patients will be “lost” (non-compliant)

    - several fractures could be prevented and then avoided.

  - All above expected benefits can lead to:

    - Decreased fracture incidence

    - Increased primary and secondary fracture prevention

    - Decreased number of people who are dependent after fracture/s with limited activities of daily living (ADL)

    - Reduction of direct and indirect cost burden on health care systems since the main cost drivers are fracture-related costs as well as indirect costs including disability and loss of productivity.

HOW WOULD IMPLEMENTING THE NOVEL REMS TECHNOLOGY SUPPORT AND FULLFILL THE ISCD MISSION INCLUDING:

• Related to, and supportive of, the ISCD mission. 
• Sufficient scientific research exists to warrant the creation of a task force and convening of a PDC, such that there is a reasonable probability that scientifically supported Position Statements can be established.
• Sufficient controversy or uncertainty in the field around aspects of the topic such that Position Statements developed by a PDC would contribute to the field.
• Impact of the official positions needs to be applicable to the constituencies ISCD serves. The greatest emphasis should be on clinical application/clinical utility.
• Topics should be supportive of, and help direct, ISCD educational programs.

- primary ISCD mission is to “advance high quality musculoskeletal health assessment in the service of superior patient care.”

Discussion - Radiofrequency Echographic Multi Spectrometry (REMS) is a quantitative ultrasound technology that has been developed to measure, assess and report specific and clinically significant characteristics of human bone in a manner that complies with the WHO established criteria for the diagnosis of osteoporosis (1). REMS has shown the capacity to provide a reliable (accurate) and reproducible (precise) measurement of BMD as compared to DXA in multiple clinical studies (12-17). In the cited studies, multiple authors converged on the conclusion that there is a very high correlation between DXA-measured and REMS-measured BMD and T-score values in the assessed populations, and that REMS had high accuracy, sensitivity and specificity for the diagnosis of osteoporosis compared to DXA. REMS is the first clinically available method for the direct measurement of lumbar and femoral bone mineral density (BMD) that does not use ionizing radiation.  REMS rapidly generates a medical report that includes BMD, T-score and Z-score. In addition, REMS automatically assess the quality of bone micro architecture, independent of BMD, and provides a five-year probability of a major osteoporosis fracture or hip fractures. REMS advances the quality of musculoskeletal health assessment and care.

- secondary mission of the ISCD is “to make quality musculoskeletal health assessment accessible for people everywhere.”

Discussion - The absence of ionizing radiation and the portable nature of the REMS technology (EchoS) allows for osteoporosis diagnosis at the point of care more broadly and at an earlier stage. It also enables accurate monitoring post diagnosis.  These features as well as others such as ease of use, accurate and reproducible results and the possibility of bed-side scanning, when combined with economic advantages of less costly system and the cost savings from earlier detection and treatment of osteoporosis, help expand access to quality musculoskeletal health care and fracture risk reduction for people everywhere.

-sufficient scientific research exists to warrant the creation of a task force and convening of a PDC, such that there is a reasonable probability that scientifically supported Position Statements can be established:

Discussion – A plethora of scientific articles and abstracts supporting the use of REMS technology exists as is demonstrated by references (7-17, 1, 20, 23, 27, 33, 39-41, 43-45, 53-55) 1. REMS has shown the capacity to provide a reliable (accurate) and reproducible (precise) measurement of BMD as compared to DXA in multiple clinical studies (12-17, 53). A cross-sectional clinical study examined the ability of the REMS derived Fracture Score (FS) to distinguish between fractured and non-fractured patients (10). The FS has shown the ability to discriminate between patients who experienced an osteoporotic fracture from those who did not (sensitivity = 76%; specificity = 68%) as compared to DXA BMD (73% and 66% respectively) (10).

When considering the area under the curve (AUC) of DXA and REMS spinal or femoral T-scores, and FS, in order to evaluate the ability to identify patients with previous osteoporotic fractures from those with no fracture, FS was shown to be an effective predictor for fractures (33, 54, 55). The AUCs obtained by the FS in the discrimination between fractured and non-fractured patients were significantly higher than those obtained using T-score values (AUC=0.818 for FS, 0.655 for DXA T-score and 0.671 for REMS T-score values, with p-value <0.0001 and 0.0003). Overall, the FS is an effective predictor of 5-year fracture risk.

In addition, REMS has been successfully used in measuring changes in BMD in a study of pregnant women completed at the unit of Obstetrics and Gynecology, University of Parma, from February to December 2020 (40). A preliminary study performed in collaboration between the Institute of Clinical Physiology and the departments of Rheumatology and Oncology at two hospitals in Lecce, Italy, concluded that the resolution capabilities of the REMS technology allowed the statistically significant and clinically meaningful serial BMD measurements of the femoral neck at a period of six months in a group of patients on Denosumab (39, 43, 44).

- applicable to the constituencies ISCD serves. The greatest emphasis should be on clinical application/clinical utility:

Discussion – REMS is directly applicable to the constituencies ISCD serves; clinicians involved in bone health assessment. The EchoS ultrasound system is a fantastic clinical tool. With this clinical tool the current users have been able to create and build solid Bone Health Programs that can serve as the Fracture Liaison Service (FLS) for hospitals and surrounding communities. The ability to provide in office examinations and to interact with patients who can watch as their bones are being assessed in real-time with the opportunity to review results immediately upon completion of the examination, is a significant attribute over DXA generated reports sent over from a local imaging center. Also, since REMS is a non-ionizing and safe diagnostic approach, there are no OSHA or FDA compliance issues. The EchoS is portable, which provides the opportunity to not only utilizing the same unit in multiple offices, but also to be able to offer bone health services to individuals who have mobility issues.

- supportive of, and help direct, ISCD educational programs:

Discussion – Radiofrequency Echographic Multi Spectrometry (REMS) is a quantitative ultrasound technology that has been developed to measure, assess and report specific and clinically significant characteristics of human bone in a manner that complies with the WHO established criteria for the diagnosis of osteoporosis (1). The development of REMS made possible to sonographically assess the lumbar spine and proximal femurs – anatomical regions that up to now, could only be measured by dual-energy x-ray absorptiometry (DXA). In addition, innovative and proprietary software development has allowed for the acquisition and automatic analysis of backscattered ultrasound waves, not only to generate B-mode images for clinical use during the REMS assessment, but also to assess bone density as well as bone quality properties. The results are rapidly reported in real-time and therefore are immediately available to both the examiner and patient.

REMS can provide the back-bone of a bone health educational program directed at current and future bone health care providers. It utilizes the same WHO designated criteria for the diagnosis of osteoporosis as DXA does, but because of the technical marvels of the EchoS system there is no need to dedicate long time for the training of the operators as it is the case of DXA, in order to prevent typical pitfalls of scanning and reporting. Instead, timesaving and clinical-effectiveness of REMS can result in employing more time to teaching and understanding the pathophysiology of osteoporosis and how to correctly treat it including the nutritional, exercise and pharmaceutical treatment modalities.

REMS is a portable tool that can be used by densitometrists and other bone health clinicians to spread the word of “Bone Health” around the country and around the world. It is not confined to a specialized room in an imaging center locked in by OSHA requirements and the fear of ionizing radiation. REMS is free to travel all roads to those in need of good bone health care. With its ease of use and immediate accurate test results, health care providers can educate their patients on the need to maintain healthy bones to prevent a life-altering and often life-threatening fragility fractures.

REFERENCES:

- WHO Scientific Group on the Assessment of Osteoporosis at Primary Health Care Level, Summary Meeting Report, Brussels, Belgium, May 2004.

- Didier Hans, et al, Quantitative ultrasound (QUS) in the management of osteoporosis and assessment of fracture risk, Journal of Clinical Densitometry, (2017) 20: 322-333.

- Marc-Antoine Krieg, et al, Quantitative ultrasound in the management of osteoporosis: The 2007 ISCD official positions, Journal of Clinical Densitometry (2008) 11: 163-187.

- Aldolfo Diez-Perez, et al, Prediction of absolute risk of non-spinal fractures using clinical risk factors and heel quantitative ultrasound, Osteoporosis Int (2007) 18: 629-639.

- Margaret A. Paggiosi, et al, A European comparison of quantitative ultrasound measurement variables: the OPUS study, Osteoporosis Int (2012) 23:2815-2828.

- M. Y. Chen, et al, Quantitative ultrasound and fracture risk prediction in non-osteoporotic men and women as defined by WHO criteria, Osteoporosis Int (2013) 24: 1015-1022.

- Francesco Conversano, et al, A novel ultrasound methodology for estimating spine mineral density, Ultrasound Med Biol, (2015) 41: 281-300

- Sergio Casciaro, et al, An advanced quantitative echosound methodology for femoral neck densitometry, Ultrasound Med Biol (2016) 42:1337–1356.

- Marco Peccarisi, et al, In-vitro study of human proximal femur microstructure: analysis of the relationship between micro-computed tomography data and quantitative ultrasound parameters, IET Science, Measurement & Technology, (2016) 10: 193-199.

- Paola Pisani, et al, A quantitative ultrasound approach to estimate bone fragility: A first comparison with dual X-ray absorptiometry, Measurement (2017) 101: 243-249.

- Aldolfo Diez-Perez, et al, Radiofrequency Echographic Multispectrometry for the in-vivo assessment of bone strength: state of the art – outcomes of an expert consensus meeting organized by the European Society for Clinical and Economic Aspects of Osteoporosis, Osteoarthritis and Musculoskeletal Diseases (ESCEO), Aging Clinical and Experimental Research (2019) 31: 1375-1389.

- Marco Di Paola, et al, Radiofrequency Echographic Multispectrometry compared with dual X-ray absorptiometry for osteoporosis diagnosis on lumbar spine and femoral neck, Osteoporosis International, (2019) 30(2):391-402.

- Bernard Cortet, et al, Radiofrequency Echographic Multispectrometry (REMS) for the diagnosis of osteoporosis in a European multicenter clinical context, Bone, (2021) 143:115786.

- Debora Meira Ramos Amorim, et al, New technology REMS for bone evaluation compared to DXA in adult women for the osteoporosis diagnosis: a real-life experience, Archives of Osteoporosis (2021) 16: 175, 1-8.

- Maria Dea Tomai Pitinca, et al, Could Radiofrequency Echographic Multi Spectrometry (REMS) overcome the limitations of BMD by DXA related to artifacts? A series of 3 cases, Journal Ultrasound Med (2021), 9999: 1-5.

- Carla Caffarelli, et al, Ability of Radiofrequency Echographic Multi Spectrometry to identify osteoporosis status in elderly women with type 2 diabetes, Aging Clinical and Experimental Research, (2021) 34(1):121-127.

- Giovanni Adami, et al, Radiofrequency Echographic Multis Spectrometry for the prediction of incident fragility fractures: A 5-year follow-up study, Bone, (2020) 134:115297.

- DHHS - FDA, Indications for Use, Echolight S.p.a 510(k) Premarket notification, Feb 15, 2018.

- Anna Nowakowska-Płaza, et al, Diagnostic agreement between radiofrequency echographic multispectrometry and dual-energy X-ray absorptiometry in the assessment of osteoporosis in a Polish group of patients, Pol Arch Intern Med (2021) 131(9):840-847.

- National Guidance for diagnosis, risk stratification and continuity of care following fragility fractures, Ministerial Guidelines 2021.

- ISCD Clinician Course, Lecture 4 - Monitoring BMD Change, US Version, December 2020.

- ISCD Clinician Course, Lecture 6 - Can you trust this DXA report, US Version, December 2020.

- Loredana Cavalli, The REMS technique is not affected by arthrosis artifact, which can hinder the densitometric recognition of osteoporosis, Ann Rheum Dis - Scientific Abstracts, 15 June 2019, 1361.

- Association of quantitative heel ultrasound with history of osteoporotic fractures in elderly men: the ESOPO study, Osteoporosis Int (2005) 16: 1749-1754.

- Adele L. Boskey, Using bone quality to assess fracture risk, AAOS Now, American Academy of Orthopedic Surgery, July 1, 2011.

- Eleftherios Paschalis, R. Mendelsohn, et al, Infrared Assessment of bone quality – a review, Clin Orthop Relat Res (2011) 469: 2170-2178.

- Tommaso De Marco, et al, A new approach for measuring the trabecular bone density through the echo sound backscattering: An ex vivo validation on human femoral heads, Measurement, (2016) 87: 51-81.

- Fjola Johannesdottir, et al, Fracture prediction by computed tomography and finite element analysis: current and future perspectives, Curr Osteoporos Rep (2018) 16:411.

- AM Cheung, et al, High-resolution peripheral quantitative computed tomography for the assessment of bone strength and structure: a review by the Canadian Bone Strength Working Group, Curr Osteoporos Rep (2013) 11:136–146

- James F Griffith, et al, New advances in imaging osteoporosis and its complications, Endocrine (2012) 42:39–51.

- Barbara C Silva, et al, Trabecular bone score: a noninvasive analytical method based upon the DXA image, J Bone Miner Res (2014) 29:518–530.

- Aldolfo Diez-Perez, et al, Treatment failure in osteoporosis, Osteoporosis Int, (2012) 23: 2769-2774.

- Kimberly Zambito, Fiorella Anna Lombardi, et al, REMS technology for the assessment of fracture risk through Fragility Score evaluation, abstract poster presentation, Orthopaedic Research Society, Tampa FL, February 2022.

- Gillespie, C. and Moran, P. Trends and disparities in osteoporosis screening in women in the United States, 2008-2014.  2017.  American Journal of Medicine: 130:  305-317.

- U.S. Department of Health and Human Services.  National Healthcare Disparities Report 2011. Rockville, MD: Agency for Healthcare Research and Quality; 2012.

- U.S. Preventive Task Force. Screening for osteoporosis:  U.S. preventive services task force recommendation statement. Ann Intern Med. 2011: 154(5): 356-364.

- U.S. Preventive Task Force.  Screening for osteoporosis to prevent fractures:  U.S. preventive services task force.  JAMA. 2018: 313(24): 2521-2531.

- Barbara C. Silva, et al, Trabecular Bone Score: a non-invasive analytical method based on the DXA image, Journal of Bone and Mineral Research, (2014), 29:518-530

- Eugenio Quarta, et al, Short-term monitoring of Denosumab effect in breast cancer patients receiving aromatase inhibitors using REMS technology on the lumbar spine.

- Anna Degennaro Valentina, et al, First assessment of bone mineral density in healthy pregnant women by means of Radiofrequency Echographic Multispectrometry (REMS) technology, European Journal of Obstetrics and Gynecology and Reproductive Biology, (2021) 263: 44-49.

- E. Kirilova, et al, Radiofrequency Echographic Multispectrometry for the assessment of both femora in pregnant women, Scientific Abstracts, Ann Rheum Dis., (2021).

- Carmelo Messina, et al, Prevalence and type of errors in dual-energy x-ray absorptiometry, European radiology, (2015) 25:1504-11

- Forcignanò R., Lombardi F.A., Ciccarese M., Quarta E., Grimaldi A., Quarta L., Ciardo D., Casciaro E., Pisani P., Conversano F., Muratore M., Casciaro S. “Short-Term Monitoring of Denosumab Effect in Breast Cancer Patients Receiving Aromatase Inhibitors Using Radiofrequency Echographic Multi-Spectrometry (REMS) Technology on Femoral Neck.” ACR Convergence 2020, November 5–9.

- Ciardo D., et al, REMS technology for short-term monitoring of Denosumab therapeutic effect in breast cancer patients receiving Aromatase Inhibitors based therapy, Virtual World Congress on Osteoporosis, Osteoarthritis and Musculoskeletal Diseases (WCO-IOF-ESCEO) 2020, 20–22 August.

- Sergio Casciaro, New perspectives in echographic diagnosis on hip and spine, Clinical Cases in Mineral and Bone Metabolism (2015), 12; 142-150.

- Olivia c. Lee, et al, The impact of diabetes on osteoporosis management and secondary fracture risk after primary fragility fractures: a propensity score-matched cohort study, JAAOS, (2022) 30:204-212.

- Giorgia Grassi, et al, Radiofrequency Echographic Multi Spectrometry (REMS) evaluation in patients with primary osteoporosis and primary hyperparathyroidism, Endocrine Abstracts  Bone and Calcium, (2020) 70 AEP134.

- El Maghraoui, et al, Reproducibility of bone mineral density measurements using dual X-ray absorptiometry in daily clinical practice, Osteoporosis Int. (2005) 16: 17-42-1748

- M,C Lodder, et al, Reproducibility of bone mineral density measurement in daily practice, Ann Rheum Dis, (2004) 63:285-289

- P. Ravaud, et al, Individual smallest detectable difference in bone mineral density measurements, Journal of Bone and Mineral Research, (1999) 14: 1449-1456

- Quality Check Module V.1.0, User Manual, April 2020, Echolight S.p.A.

- EchoStudio V.2.0, User Manual, September 2020, Echolight S.p.A.

- Pitinca, et al, Use of REMS technology in patients with spine artifacts: A new diagnostic opportunity, International Society for Clinical Densitometry Congress, ISCD, 2020

- Ciardo D., et al, Incident fracture risk prediction using the Fragility Score Calculated by lumbar spine Radiofrequency Echographic Multi Spectrometry (REMS) scans, European E-Congress of Rheumatology (EULAR), 2-5 June 2021

- Ciardo D., et al, The effectiveness of Fragility Score measured by Radiofrequency Echographic Multi Spectrometry (REMS) for the assessment of incident fracture risk, World Congress on Osteoporosis, Osteoarthritis and Musculoskeletal Diseases (WCO-IOF-ESCEO), 26–29 August 2021.

APPENDIX I

CLINICAL USERS - LETTERS IN SUPPORT OF THE ECHOS SYSTEM AND REMS TECHNOLOGY IN THE BONE HEALTH PROGRAMS AND THE TREATMENT OF OSTEOPOROSIS AND THE PREVENTION OF FRACTURES

Andrew P. Bush, MD, FAAOS, CWSP

Central Carolina Orthopaedic Associates, PLLC

1139 Carthage Street, Suite 101

Sanford, NC 27330

919.774.1355

www.centralcarolinaortho.com

Curtis Foy, MD, CCD

Rheumatology Associates of Delmarva

505 Dutchman’s Lane, Suite A3

Easton, MD 21601

Phone 410-819-6545

Fax 410-819-6750

Yevgeniya Kushchayeva, MD, PhD

USF Diabetes & Endocrinology Center

UNIVERSITY OF SOUTH FLORIDA

12901 Bruce B. Downs Blvd, MDC 62

Tampa, FL  33612

Tel: (813) 821-8011

Fax: (813) 905-8970

kushchayevay@usf.edu

Edmund J. MacLaughlin, M.D., FACR, CCD

Rheumatology Associates of Delmarva

505 Dutchman’s Lane, Suite A3

Easton, MD 21601

Phone 410-819-6545

Fax 410-819-6750

Kimberly Zambito, M.D., FAAOS, FAOA

Mercer Bucks Orthopaedics

Hamilton Office

2501 Kuser Road, Suite 3

Hamilton, NJ 08691

855-896-0444

www.mbortho.com

February 2, 2022

###### ISCD Headquarters

955 South Main St. B202
Middletown, CT 06457 USA

P: 1-860-259-1000
E: iscd@iscd.org

Christopher R. Shuhart, MD, MHA, CCD

Position Development Conference Steering Committee

Dear Dr. Shuhart:

As an ABOS certified orthopedic surgeon practicing for over 25 years, I have surgically treated a multitude of patients who sustained low-energy pathological (fragility) fractures due to poor bone health. It was only recently however, that I realized surgically fixing a fracture with titanium or stainless steel, although absolutely necessary, was in reality only a partial treatment and at best a temporary fix – there was a much more significant issue that needed to be addressed to really treat these patients. The underlying reason why bones fracture in the first place absolutely needs to be determined and addressed not only to prevent additional fractures from occurring but to keep the patient from ever sustaining their first fracture. This is an entirely new way of practicing medicine for a traditionally-trained orthopedic surgeon and there are many aspects to the field of Bone Health Medicine which were not taught in our training program.

The ISCD was one of the organizations that was highly recommended to me as a resource for physicians and other health care providers interested in providing appropriate bone health care to their patients. Participating in ISCD courses and utilizing ISCD educational resources has provided me with the foundation of knowledge that was necessary to expand the scope of my orthopedic practice and to build a Bone Health Program to provide the best care that I can provide to my patients.

In developing my Bone Health Program, I utilized resources available to me as a member of the ISCD. I have found the ISCD Position Papers to be valuable - with the current plethora of available technologies and treatment modalities, it is of significant benefit to have structured and well-researched guidelines to help physicians provide their patients with the best advice and safest methods of treatment. For this reason, I am submitting my request to the PDC to review and consider issuing a Position Statement on the use of Radiofrequency Echographic Multi Spectrometry (REMS) as a clinically appropriate and useful tool to assess bone health and fracture risk. It is a tool that I have been successfully using in my practice and has allowed me to provide a level of care that otherwise I would have not been able to deliver to my patients. Therefore, it is my goal with this request to the PDC, to provide the reviewers with supporting academically-sound data as well as my personal letter in support of this technology and additional letters in support of REMS from several of my colleagues who are currently using the EchoS system in their clinical practices.

Thank you for your consideration. If I can be of further assistance, please do not hesitate to contact me.

Sincerely,

Andrew P. Bush, MD, FAAOS, CWSP

Diplomat, American Board of Orthopedic Surgery

Central Carolina Orthopaedic Associates, PLLC ● 1139 Carthage Street, Suite 101 ● Sanford, NC 27330 ● 919.774.1355

Edmund J. MacLaughlin, MD, FACR, CCD

Elizabeth S. Clayton, MD, FACR

Naila B. Ahmad, MD, FACR

Curtis M. Foy, MD, CCD

Catherine G. Campbell, PA-C

Meghan B. McCarter, FNP

505 Dutchman’s Lane, Suite A3  	                  			10 Aurora Street				1415 S. Division Street

Easton, MD 21601			  		 	 Cambridge, MD 21613			Salisbury, MD 21804

Phone 410-819-6545			  		 	 410-228-0556				410-860-6801

Fax 410-819-6750			   			 410-228-3986				410-860-6864

ISCD Headquarters

955 South Main St. B202

Middletown, CT 06457

P: 1-860-259-1000

E: iscd@iscd.org

Christopher R. Shuhart, MD, MHA, CCD

Position Development Conference Steering Committee

Dr. Shuhart:

My name is Dr. Curtis Foy.  After attending college and medical school in New Jersey, I completed a family medicine residency through the United States Air Force and have been board certified in family medicine since 2002.  In 2010, after eleven years of active-duty service, I moved to the Eastern Shore of Maryland and joined a private practice primary care group which was ultimately absorbed by a hospital system.

In May of 2021, I was offered a unique opportunity to join a local rheumatology practice to help expand their osteoporosis and bone health programs.  The owner of the practice, Dr. Edmund MacLaughlin, has been a board-certified rheumatologist for decades and is one of the earliest adopters of all things related to bone health.  He was the first physician on the Eastern Shore to own an office based DXA scanner, and bone health has truly been one of his passions over the decades. After joining this practice and enveloping myself in the world of osteoporosis, I realized just how much I had been missing with these patients. The “front line” physicians and providers in family medicine, internal medicine, and gynecology are doing their best to keep up with the crush of patients and their needs.  I was one of those providers, strapped for time and woefully undereducated on DXA interpretation, what medications to use when, etc.  Suddenly, DXA interpretation errors from outside imaging centers stuck out like a sore thumb.

I learned about the REMS ultrasound system shortly after I arrived at my new job last year, and the more I read about it the more intrigued I became about its potential role in initial assessment and follow-up for patients with osteoporosis.  One of the things that I was especially excited about is the portability of the system.  Medical practices that incorporate bone health with multiple locations and can’t afford to put a DXA machine at every location could benefit from the portability of a REMS system.  Nursing home and ALF patients are an especially vulnerable population to fracture AND many times there are access to care issues for them, creating a clinical paradox in which the population most vulnerable to fragility fracture are the most challenging to assess appropriately.  Being able to take a REMS system into those types of facilities for patient assessments would be a huge step toward eliminating some of the disparities in access to care.  A nursing home or ALF could even purchase one of these machines themselves.  There is no need for a dedicated room to house equipment.

Thousands of DXA scanners have disappeared from physician’s offices over the years as reimbursements have declined, and an ever-increasing percentage of these scans are now done in large radiology centers.  Very few of these centers have ISCD certified technicians or radiologists.   At the hospital system where I last worked, the mammography technician would pull ‘double duty’ on the DXA scanner.  Even though I have only been doing this for less than a year, I’ve caught more errors on these scans than I can count.  In preparation for my new role, I was loaned (and later bought my own) a copy of Dr. Sydney Bonnick’s excellent book “Bone Densitometry in Clinical Practice” which points out many of the problems that can arise with DXA image acquisition and interpretation.  Even though the third and most recent edition of the book is 13 years old, imaging centers currently doing DXA have not gotten the memo and continue to make the same errors.  Unless the physician or provider ordering the study understands how to thoroughly vet it, they may end up making clinical decisions based on inaccurate data.  Speaking anecdotally, I can tell you that my ability to properly vet a DXA scan as a primary care provider was extremely limited, and I feel like I was one of the better physicians in terms of trying to keep up with everything!

There is also the issue of consistency between bone density assessments on different machines.  We live in an increasingly mobile society and for a few of my patients I have DXA results spanning a decade that have been done in three or four different facilities across the country.  All different machines and models, dissimilar positioning, etc.  Some inferences of course can be made about the general direction of bone density, but the error of the machine, the error of the technician, and the differences between machines must all be considered.   A REMS machine’s margin of error, on the other hand, is much lower even when considering different technicians doing the same patient.  A REMS machine in Miami would theoretically be very close to a REMS machine in Washington DC.   Decision making about treatment would, in turn, be able to be made with greater confidence.

As our population continues to age, new solutions will be required to minimize the damage and cost that fragility fractures will inflict on this population.  While DXA scanners are amazing tools in the right hands, we also need to consider their limitations and look at alternate ways that accurately capture fracture risk.  Reducing health care disparities and improving access to care with a portable solution seems to be a good direction to move towards.

Sincerely,

Curtis Foy, MD, CCD

Rheumatology Associates of Delmarva

Mobile:  410-226-6420

E-mail:  foycur@gmail.com

USF DIABETES and ENDOCRINOLOGY CENTER

UNIVERSITY OF SOUTH FLORIDA

February 5, 2022

To:  ISCD Headquarters

955 South Main St. B202
Middletown, CT 06457 USA

P: 1-860-259-1000
E: iscd@iscd.org

From: Yevgeniya Kushchayeva, MD, PhD

Dear Dr. Shuhart:

I am a board-certified endocrinologist with a particular interest in bone health. I am building the Osteoporosis Program at the University of South Florida. ISCD recommendations are critical very valuable and for all of us, practicing providers focused on bone health since ISCD is a leading organization in providing guidance in osteoporosis diagnosis, management, and bone-related updates.

Recently I had a chance to use a new FDA approved diagnostic modality for bone health assessment - Radiofrequency Echographic Multi Spectrometry (REMS). I applied this technique in situations where DXA use was limited, such as severe skeletal deformities and presence of instrumentation. The data I have received helped me to guide treatment and even helped with anti-osteoporosis medication approval by an insurance company.

Even we all use the gold standard for osteoporosis diagnosis that is DXA, sometimes the coverage of bone-targeted medications by insurance companies depends on parameters measured during the DXA evaluation. In situations when we cannot receive convincing data with DXA, REMS might be an good alternative to assess bone mineral density and advocate for appropriate treatment.

Since REMS has already received the clearance of the U.S.A. Food and Drug Administration for the measurement of the diagnostic parameters BMD, T-score and Z-score, and for monitoring bone changes in the clinical routine and has been validated through several national studies each time focused on specific age ranges, I was wondering if you might consider adding REMS as a recommended tool for bone assessment of patients to the ISCD guidelines which we all closely follow.

I appreciate your consideration.

If you might have any questions, please do not hesitate to contact me.

Sincerely,

Yevgeniya Kushchayeva, MD, PhD

USF Diabetes & Endocrinology Center

kushchayevay@usf.edu

______________________________________________________________________________

Mailing Address:  12901 Bruce B. Downs Blvd, MDC 62, Tampa, FL  33612 • Tel: (813) 821-8011 • Fax:  (813) 905-8970

Edmund J. MacLaughlin, MD, FACR, CCD

Elizabeth S. Clayton, MD, FACR

Naila B. Ahmad, MD, FACR

Curtis M. Foy, MD, CCD

Catherine G. Campbell, PA-C

Meghan B. McCarter, FNP

505 Dutchman’s Lane, Suite A3  	                  			10 Aurora Street				1415 S. Division Street

Easton, MD 21601			  		 	 Cambridge, MD 21613			Salisbury, MD 21804

Phone 410-819-6545			  		 	 410-228-0556				410-860-6801

Fax 410-819-6750			   			 410-228-3986				410-860-6864

February 8, 2022

ISCD Headquarters

955 South Main St. B202

Middletown, CT 06457

P: 1-860-259-1000

E: iscd@iscd.org

Christopher R. Shuhart, MD, MHA, CCD

Position Development Conference Steering Committee

Dear Dr. Shuhart:

I would like to amplify the issues raised by Dr. Foy regarding the need for the Echolight REMS technology to be evaluated by Position Development Conference through ISCD.

Largely because of inadequate reimbursement for DXA, exacerbated by the high cost of operating a DXA facility (the cost of the machine, the cost of moving and calibrating the machine, the cost of an annual maintenance contract and the large amount of floor space required by the machine), well over 5,000 DXA machines have been lost through the United States in the last decade creating a serious access to care issue in many localities.

Despite decades of effort by ISCD, we have barely dented the continuing challenge of poor quality DXA studies and poor quality interpretation leading to errors in treatment.

Echolight REMS appears to have the potential to move us past both of these log jams.  In addition, we are all aware of the limitation of DXA in identifying the true fracture risk of individuals with various conditions that affect bone quality, which is not measured by DXA.

For the first time, we have a portable machine without radiation exposure that has the potential to measure both quantity and quality of the bones that are at the greatest risk for fracture, the hip and lumbar spine directly.

As this machine is now approved for use in the United States and as is already in use in large areas of Europe and elsewhere and given that the ISCD is meant to be “international” in scope, I request that this technology be included in the upcoming Position Development Conference agenda.

Sincerely,

Edmund J. MacLaughlin, M.D., FACR, CCD

EJM:jlm

Hamilton Office         Lawrenceville Office        Langhorne	 Office                         Princeton Office              Marlton Office         Monroe Office

2501 Kuser Road        2A Princess Road               1203 Langhorne-Newtown Rd   10 Forrestal Road South 651 N. Route 73      111 Union Valley Rd

Suite 3                          Suite 200                             Suite 220                                       Suite 105                            Suite 405                  Suite 102

Hamilton, NJ 08691   Lawrenceville, NJ 08648   Langhorne, PA 19047                  Princeton, NJ 08540        Marlton, NJ 08053  Monroe Township, NJ 08831

ISCD Headquarters

955 South Main St. B202

Middletown, CT 06457

P: 1-860-259-1000

E: iscd@iscd.org

Christopher R. Shuhart, MD, MHA, CCD

Position Development Conference Steering Committee

Dear Dr. Shuhart:

I am a board-certified orthopaedic surgeon with fellowship training in hand surgery.  Over the years, my colleagues and I have missed opportunities to engage patients who have sustained fragility fractures about their bone health.  While the Own the Bone Program has attempted to engage orthopaedic surgeons to address osteoporosis, we are still missing the mark.  During my career, I have treated a number of distal radius fractures in younger women.  In the past, I have educated these patients about this type of fracture as a sentinel event, then referred them to a non-orthopaedic colleague to evaluate them for osteoporosis, later to learn that the patients have fallen through the cracks.  As a result, I decided to start a comprehensive bone health program within my practice.  My focus is to engage people about their bone health prior to fracture, as well as, to work with patients to prevent future fractures in those who have already sustained a fragility fracture.  I have recognized the need to do a better job of reaching out to patients to provide them with reliable information, including providing them with opportunities to use tools other than DXA to evaluate their bone health.

I began using REMS technology to offer my patients an in-office tool to evaluate bone mineral density and microarchitecture as represented by the Fragility Score.  I was surprise to have so many women reach out to me to request an evaluation with this technology.  These women were frustrated with their DXA results and felt disheartened that they had done everything right in terms of nutrition, exercise, and lifestyle.  Many of these women were consumed by the diagnosis of osteoporosis and began to self-educate and search for something that would give them additional information about their bone health.  For them, the DXA results did not tell the complete story.  They wanted information about the microarchitecture of their bone.   In short, REMS technology has given me the opportunity to reach many women and men who would have otherwise abandoned traditional medicine to seek treatment.  For my practice, this technology has given me the opportunity to engage patients on a deeper level.  While their DXA T-score and REMS T-score may be similar, it is the additional information about microarchitecture that has allowed me to offer reassurance and hope that a particular treatment may be the right treatment for their circumstance.  I have been able to successfully use REMS technology in parallel with DXA.

For these reasons, my colleagues and I request your assistance an issuing a Position Statement on the use of Radiofrequency Echographic Multi Spectrometry (REMS) as a clinically appropriate and useful tool to assess bone health and fracture risk.  It is an impactful in-office technology which can help us reach out to patients to advance our goal of treating bone health and preventing future fractures.

Thank you for your consideration.  Please contact me if I can be of further assistance.

Respectfully,

Kimberly Zambito, M.D., FAAOS, FAOA

Diplomat, American Board of Orthopaedic Surgery

kzambito@mbortho.com

mobile: 267-240-6599

APPENDIX II

SUMMARY TABLE OF FEATURES OF REMS AND DXA AND DXA+TBS

| Summary table of Features of REMS vs. DXA and DXA + TBS |  |  |  |
| --- | --- | --- | --- |
| Attribute | REMS | DXA | DXA + TBS |
| Ability to determine Axial BMD - spine and femur | √ | √ |  |
| Ability to determine Axial BMD - spine | √ | √ | √ |
| Ability to provide T-score | √ | √ | √ |
| Ability to diagnose Osteoporosis | √ | √ | √ |
| Ability to monitor osteoporosis treatment | √ | √ | √ |
| Ability to monitor BMD in Diabetes Type 2 | √ |  | √ |
| Ability to monitor BMD in hyperparathyroidism | √ | √ | √ |
| Ability to monitor BMD in Osteogenesis Imperfecta | √ |  |  |
| Ability to monitor BMD in children | In development | Limited |  |
| Ability to monitor BMD during pregnancy | √ |  |  |
| Ability to determine BMD independent of bone size | √ |  |  |
| Ability to provide a VFA |  | √ | √ |
| Ability to exclude bone artifacts | √ |  |  |
| Affected by the “adipose affect” |  | √ | √ |
| Ability to generate an immediate report | √ |  |  |
| Ability to predict fracture risk | √ | Limited | √ |
| Ability to predict medium term fracture risk (5 year) | √ |  |  |
| Ability to predict fracture risk without FRAX | √ |  |  |
| Ability to predict fracture risk without BMD | √ |  |  |
| Ability to assess bone quality | √ |  | √ |
| Ability to be easily transported | √ |  |  |
| Energy source non-ionizing radiation | √ |  |  |
| Patient position not critical for accurate examination | √ |  |  |
| Requires OSHA and nuclear safety inspections |  | √ | √ |
| Requires routine recalibration |  | √ | √ |
| Dependent on image quality |  | √ | √ |
| Smallest Detectable Difference (SDD) | 0.009 g/cm2 | 0.030 g/cm2 | N.D. |
| Least Significant Change (LSC) | 0.38 – 1.05 % | 4.07 – 5.60 % | N.D. |
| Intra-Operator Repeatability | 0.32 – 0.36 % | 1.47 – 2.02 % | N.D. |
| Inter-Operator Repeatability | 0.48 – 0.54 % | N.D. | N.D. |
| Short term (6 month) assessment of treatment | √ |  |  |
| Determination of body composition | √ |  |  |
| Determination of sarcopenia | In development |  |  |